Technical Bulletin BTM-TB-002

July 15st, 2022

BOX TERMINAL MULTI Software Update

This document explains the procedure to install the DDL and Software 1.9.2022.5 on a Box Terminal Multi with the aim to upgrade the Software and enhance the video-traces sync and stability of device. Moreover, this installation adds a watchdog service which continuously monitors the software and turns the Box Terminal Multi application on in case of any crash.

Installation

Follow the instructions below to install the sw 1.9.2022.5 on the Box Terminal Multi:

- Copy “BTM – sw 1.9.2022.5” folder locally

- Extract the folder content

- Press Windows+R to open “Run” box. Type “cmd” and then click “OK” to open a Command Prompt

- Use “cd” to change directory and move inside the “BTM – sw 1.9.2022.5” directory

- Type: “micromedSenderToBTM.bat <ipaddress>” where <ipaddress> is the Box Terminal Multi IP address you want to update (e.g. “micromedSenderToBTM 10.1.2.128”)

- Press ENTER

- At the end of the installation, the Box Terminal Multi reboots automatically.

- After the reboot, check the software version: after pressing the four corners of the Box Terminal Multi display, the software version will appear at the top right corner