# Box Terminal Multi v 1.02.0004 – Release Test Plan

## Test Environment:

Processor: Intel(R) Core(TM) i7-8700 CPU @ 3.20GHz

RAM: 8.00 GB

Operating System: Windows 10 Pro (1903)

2 Monitor, display resolution: 1920x1080

Acquisition Software: SystemPLUS Evolution v1.06.04 and SystemPLUS Evolution v1.04.215

Acquisition Interface: BOX TERMINAL MULTI

Firmware version: 2021.1

Software Version: 1.9.2022.5

Acquisition Devices: SD LTM 32 PLUS, SD LTM 64 PLUS (x4)

Corio: C2-610A CORIOview DVI-U Multiviewer

Tests performed by Federico Bozzato, Laura Grimaldi, Antonio Marittimi

## Box Terminal Multi V1.02.0004 VALIDATION TEST PLAN

| ID | Test Suite | Result |
| --- | --- | --- |
| 342 | Box Terminal Multi fw 1.9.2022.5 Test Suite | Passed |
| 341 | Box Terminal Multi Regression Test Suite | Passed |

### 342. Box Terminal Multi fw 1.9.2022.5 Test Suite

| ID | Test Case | Test Script | Result |
| --- | --- | --- | --- |
| 1754 | VEEG no more synchronized after the recovery exceeds 2GB with blackmagic and 256 channels Test Case | VEEG no more synchronized after the recovery exceeds 2GB with blackmagic and 256 channels Test Script | Passed |
| 1755 | VEEG no more synchronized when there is the trace change after a recovery using IP Axis Camera and 256 channels Test Case | VEEG no more synchronized when there is the trace change after a recovery using IP Axis Camera and 256 channels Test Script | Passed |
| 1756 | BTM firmware crashes if an external software opens and closes sockets at the port 2000 Test Case | BTM firmware crashes if an external software opens and closes sockets at the port 2000 Test Script | Passed |

1754. VEEG no more synchronized after the recovery exceeds 2GB with blackmagic and 256 channels Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Start a VEEG acquisition in SystemPLUS Evolution with SD LTM 256 PLUS with Box Terminal Multi and Blackmagic at 2kHz as sampling rate - Unplug the SD LTM 256 PLUS after 5 min of recording - Wait for 38 min - Plug the SD LTM 256 PLUS to the box terminal multi - After the recovery, let the acquisition run for other 5 - 10 min - Stop the acquisition - Review the trace | - During the recovery, there shall be the trace change as the 1st file reaches the 2GB limit - Both the exams shall be synchronized with the video and the audio (+- 3 frames of tolerance) | Passed |
| 2 | - Repeat the previous step with: - SD LTM 64 PLUS - SD LTM 128 PLUS | - All iterations shall succeed | Passed |

1755. VEEG no more synchronized when there is the trace change after a recovery using IP Axis Camera and 256 channels Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Start a 10min - long VEEG acquisition in SystemPLUS Evolution with SD LTM 256 PLUS with Box Terminal Multi and Axis IP Camera at 2kHz as sampling rate - Unplug the SD LTM 256 PLUS after 5 min of recording - Wait for 10 min - Plug the SD LTM 256 PLUS to the box terminal multi - After the recovery, let the acquisition run for few minutes - Stop the acquisition - Review the trace | - When the headboxes return online, there shall be the trace change - Both the exams shall be synchronized with the video and the audio (+- 3 frames of tolerance) | Passed |
| 2 | - Repeat the previous step using: - SD LTM 64 PLUS - SD LTM 128 PLUS | - All iterations shall succeed | Passed |

1756. BTM firmware crashes if an external software opens and closes sockets at the port 2000 Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Start an EEG acquisition with SystemPLUS Evolution using Box Terminal Multi and SD LTM 64 PLUS - Download the attached python script - Edit it writing the IP Address of the Box Terminal Multi to "attack" - While the acquisition is running, run the attached script from command window - Wait for the launch of an exception by the python script (it may be necessary to launch multiple instances of the script) | - During the "attack", the acquisition window popup shall display the headbox as offline - Once the python script stops, the box terminal multi shall automatically reboot. Afterwards, the acquisition shall resume and the data acquired during the "attack" shall be recovered | Passed |

### 341. Box Terminal Multi Regression Test Suite

| ID | Test Case | Test Script | Result |
| --- | --- | --- | --- |
| 1742 | Box Terminal Multi Regression - Connection Test Case | Box Terminal Multi Regression - Connection Test Script | Passed |
| 1743 | Box Terminal Multi Regression - Display Test Case | Box Terminal Multi Regression - Display Test Script | Passed |
| 1744 | Box Terminal Multi Regression - Software Update Test Case | Box Terminal Multi Regression - Software Update Test Script | Passed |
| 1745 | Box Terminal Multi Regression - Firmware Update Test Case | Box Terminal Multi Regression - Firmware Update Test Script | Passed |
| 1746 | Box Terminal Multi Regression - Headboxes Update Test Case | Box Terminal Multi Regression - Headboxes Update Test Case | Passed |
| 1747 | Box Terminal Multi Regression - 32 Channels Acquisition Test Case | Box Terminal Multi Regression - 32 Channels Acquisition Test Script | Passed |
| 1748 | Box Terminal Multi Regression - 64 Channels Acquisition Test Case | Box Terminal Multi Regression - 64 Channels Acquisition Test Script | Passed |
| 1749 | Box Terminal Multi Regression - 128 Channels Acquisition Test Case | Box Terminal Multi Regression - 128 Channels Acquisition Test Script | Passed |
| 1750 | Box Terminal Multi Regression - 256 Channels Acquisition Test Case | Box Terminal Multi Regression - 256 Channels Acquisition Test Script | Passed |
| 1751 | Box Terminal Multi Regression - Camera Control Test Case | Box Terminal Multi Regression - Camera Control Test Script | Passed |
| 1752 | Box Terminal Multi Regression - Synchronism with IP HD Camera Test Case | Box Terminal Multi Regression - Synchronism with IP HD Camera Test Script | Passed |

1742. Box Terminal Multi Regression - Connection Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Power on BTM | - Box Terminal Multi display shall turn on and show a Micromed loading picture after few seconds. | Passed |
| 2 | - Connect Box Terminal Multi with the acquisition PC via ethernet cable. Alternatively, connect the ethernet cable to make the Box Terminal Multi visible online | - Ethernet icon on Box Terminal Multi display shall change form red to green | Passed |
| 3 | - Check the BTM IP address: - Touch the Box Terminal Multi display corners clockwise from top left to bottom left to open the setting window. | - Box Terminal Multi setting window shall open. - If Box Terminal Multi is online and in DHCP, the IP shall be assigned automatically | Passed |
| 4 | - Set the appropriate Box Terminal Multi IP Address touching up/down the arrows: it must be in the same network of PC (skip if Box Terminal Multi in online) | - The relative numbers in the field shall change. | Passed |
| 5 | - Verify the presence of FPGA firmware version (first row) and software version (second row) in the top right of setting windows - Check that versions are updated for the release |  | Passed |
| 6 | - Touch Apply. | - Setting window shall close and display shall go back to main window. | Passed |
| 7 | - From PC, ping the Box Terminal Multi | - Command prompt window shall open. - After typing the ping command and pressing Enter, four transmission completed messages shall be printed. | Passed |

1743. Box Terminal Multi Regression - Display Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - With Box Terminal Multi online and powered on, disconnect and connect again the ethernet cable. | - Ethernet icon shall switch red/green. | Passed |
| 2 | - Connect and disconnect an SD LTM PLUS with HBC LTM PLUS cable (64, 128 or 256 channels) | - SD LTM PLUS icon shall appear green painted or switch red/green. SD LTM PLUS serial number shall be displayed when connected - Whether 64, 128 channels or 256 channels configuration is used, the serial number displayed shall be that of the first amplifier | Passed |
| 3 | - Open SystemPLUS Evolution and set the Headbox as the one/ones connected to Box Terminal Multi | - PC icon shall become green and shall show PC IP address. | Passed |
| 4 | - Disconnect HBC LTM PLUS cable from Box Terminal Multi blue connector. | - SD LTM PLUS icon shall switch red. - PC icon shall switch red. | Passed |
| 5 | - ​​​​​​Re-connect HBC LTM PLUS cable to BOX TERMINAL MULTI blue connector. | - SD LTM PLUS icon shall appear green painted. SD LTM PLUS serial number shall be displayed when connected. - PC icon shall become green and shall show PC IP address when communication takes place. | Passed |
| 6 | - Stop the acquisition. | - PC icon shall switch red. | Passed |
| 7 | - For firmware versions higher than 2019.02 and software versions higher than 1.7.2019.4, check display backlighting management. | - Box Terminal Multi main window shall blank in 10 seconds and then shall immediately turn off after stopping the acquisition. | Passed |
| 8 | - When display is turned off, touch it to turning it on. | - Display shall turn on and main window shall appear. | Passed |
| 9 | - Enter in settings window and wait a minute. | - Display shall not blank and turn off. | Passed |

1744. Box Terminal Multi Regression - Software Update Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Run PuTTY.exe in your PC. | - PuTTY configuration window shall open | Passed |
| 2 | - Set Box Terminal Multi IP address into Host Name (or IP address) box and select SSH connection type (PuTTY_configuration) |  | Passed |
| 3 | - Click on Open. | - PuTTY terminal shall open. | Passed |
| 4 | - Type root and press Enter (PuTTY_terminal_1) | - Box Terminal Multi command line shall be shown: - root@micromedBoxTerminalMulti:~# | Passed |
| 5 | - ​​​​​​ Type the command: - # ps -A - in PuTTY terminal to see the list of running processes. | - A long list of running processes shall be printed in PuTTY terminal (PuTTY_terminal_2) | Passed |
| 6 | - Find MicromedBoxTerm process and take note of its PID number. |  | Passed |
| 7 | - Type the command - # kill MicromedBoxTerminalPID in PuTTY terminal (PuTTY_terminal_3) - Press Enter to close the running program. | - Box Terminal Multi software shall close. - Display shall be blank. When display is touched, nothing shall happen | Passed |
| 8 | - Open Windows Command Prompt and type Pscp BTM_SoftwarePath root@BTM_IPaddress:/usr to copy the new software version to BOX TERMINAL MULTI usr folder using SSH remote communication protocol. Press Enter to start the copy process. | - The copy process shall be displayed (PuTTY_terminal_4) | Passed |
| 9 | - Once the copy is finished return to PuTTY terminal, enter root/usr folder - # cd /usr - and type - # ls -l - to check if Box Terminal Multi software file has been correctly copied. | - The usr folder contents should be display, including MicromedBoxTerminal linux executable file (PuTTY_terminal_5) | Passed |
| 10 | - Type reboot command from PuTTY terminal. | - Box Terminal Multi shall restart and new software shall start up. - If the booting process freezes, switch off and switch on the power button. | Passed |
| 11 | - Enter Box Terminal Multi settings windows. Check on Box Terminal Multi display (top right) if the software version (second row) has been correctly updated. | - Software version shall be changed | Passed |

1745. Box Terminal Multi Regression - Firmware Update Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Open Windows Command Prompt and type Pscp BTM_FirmwarePath root@BTM_IPaddress:/home/root to copy the new FPGA firmware to Box Terminal Multi /home/root folder using SSH remote communication protocol. Press Enter to start the copy process | - The copy process shall be displayed (cmd_firmware) | Passed |
| 2 | - Run PuTTY.exe in your PC | - PuTTY configuration window shall open. | Passed |
| 3 | - Set BOX TERMINAL MULTI IP address into Host Name (or IP address) box and select SSH connection type (PuTTY_configuration_2) |  | Passed |
| 4 | - Click on Open | - PuTTY terminal shall open. | Passed |
| 5 | - Type root to login into Box Terminal Multi OS and press Enter. | - Box Terminal Multi command line shall be shown: - root@micromedBoxTerminalMulti:~# | Passed |
| 6 | - Type # ls to check if Box Terminal Multi firmware file micromed.bin has been correctly copied. | - The /home/root folder contents shall be display (PuTTY_terminal_6) | Passed |
| 7 | - Launch the upgradeFPGA.sh script typing: # ./upgradeFPGA.sh and wait until it finishes. | - Updating process shall be displayed (PuTTY_terminal_7) | Passed |
| 8 | - As the script requests to do, reboot the BTM toggling twice the power switch. | - Box Terminal Multi shall restart. | Passed |
| 9 | - Enter Box Terminal Multi settings windows. Check on Box Terminal Multi display (top right) if the firmware version (first row) has been correctly updated. | - Firmware version shall be changed. | Passed |

1746. Box Terminal Multi Regression - Headboxes Update Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Save the new SD LTM PLUS firmware "sdltmplus.dat" file in C:\System98\Programs . - If the headbox is already updated at the latest version, carry out a downgrade and then update again at the latest version. |  | Passed |
| 2 | - Connect the SD LTM PLUS to Box Terminal Multi with HBC LTM PLUS cable and TERM PLUS adapter |  | Passed |
| 3 | - In SystemPLUS Evolution open Tools > Options > EEG - Set the headbox to be used with Box Terminal Multi: set BqBroadcast and click on Set Device and select the correct SD LTM PLUS device. - Check that the firmware version to update is correct and click on Start Upgrade. | - The software shall show a message of correct update and on the Headbox display you shall see a countdown until 0. - The Headbox shall reboot automatically (click then on OK in SystemPLUS Evolution) | Passed |
| 4 | - Verify the correct firmware upgrade. | - When the headbox is switched on, the latest firmware version shall be shown on the display. | Passed |

1747. Box Terminal Multi Regression - 32 Channels Acquisition Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Connect Box Terminal Multi and SD LTM 32 PLUS with HBC LTM PLUS cable. | - Box Terminal Multi display shall turn on and the headbox icon shall become green. | Passed |
| 2 | - Connect waveform generator to SD LTM 32 PLUS: - Connect white MMLTM jackbox to SD LTM 32 PLUS. - Connect MMLTM jackbox to MMADP with appropriate cables into related connectors. - MMDAP must have the switch in 1/1000 mode. |  | Passed |
| 3 | - Set in the waveform generator a sinusoidal wave Vpp=100mV f=10Hz. |  | Passed |
| 4 | - Open SystemPLUS Evolution and start an EEG recording using 32 channels and 256Hz sampling rate. | - System Plus Evolution shall display the same sinusoidal waveform in each channel. | Passed |
| 5 | - Disconnect SD LTM 32 PLUS and connect it again after a while to check the correct data recovery. | - The displayed traces shall freeze until, after few second, the recovery starts | Passed |
| 6 | - Switch Box Terminal Multi off during acquisition. - Check the correct data recovery after turning Box Terminal Multi on again. | - The displayed traces shall freeze until, after Box Terminal Multi boot, the recovery starts. | Passed |
| 7 | - Stop the acquisition. | - Acquisition shall stop. - On Box Terminal Multi display, PC icon shall switch red. | Passed |
| 8 | - Start EEG recording using 32 channels and 16384Hz sampling rate | - SystemPLUS Evolution shall display the same sinusoidal waveform in each channel. | Passed |
| 9 | - Disconnect SD LTM 32 PLUS and connect it again to check the correct data recovery. | - The displayed traces shall freeze until the recovery starts. | Passed |
| 10 | - Switch Box Terminal Multi off during acquisition. - Check the correct data recovery after turning Box Terminal Multi on again | - The displayed traces shall freeze until, after Box Terminal Multi boots, the recovery starts. | Passed |
| 11 | - Stop the acquisition. | - Acquisition shall stop. - On Box Terminal Multi display, PC icon shall switch red. | Passed |
| 12 | - Check the correct working of triggers IN: - In System Plus Evolution set triggers: Tools > Settings > Trigger Settings - Set: - Trigger in on channel Fp1 - Trigger in 2 on channel Fpz - Trigger OUT on channel Fp2 - Click on Exit - Keep only Fp1, Fpz, Fp2 and MKR+ data traces. - With a RCA cable, connect GEN32 trigger out RCA connector to BOX TERMINAL MULTI trigger in 1 channel (Black RCA connector). - Switch on GEN32 - Start Acquisition at 16384Hz (See Trigger_Settings attached) | - Trigger in 1 channel Fp1 and MKR+ channel shall display triggers at the frequency set in GEN32. - Triggers shall not display in other channels. | Passed |
| 13 | - Set time base 1 second to check if multiple triggers occur. - Keep this setting for next trigger check steps. | - Single trigger signal shall be displayed. | Passed |
| 14 | - Check trigger in 2 channels: - Disconnect RCA cable from Box Terminal Multi trigger in 1 channel (Black RCA connector) and connect it to Box Terminal Multi trigger in 2 channels (White RCA connector). | - Trigger in 2 channel Fpz and MKR+ channel shall display triggers at the frequency set in GEN32. - Triggers shall not display in other channels. | Passed |
| 15 | - Check if multiple triggers occur. | - Single trigger signal should be displayed. | Passed |
| 16 | - Verify the correct working of triggers OUT: - Connect GEN32 trigger out RCA connector to Box Terminal Multi trigger out channel (Red RCA connector). | - Trigger Out Channel shall not display any triggers. | Passed |
| 17 | - Verify the correct working of triggers OUT: - With RCA cable, connect Box Terminal Multi Trigger Out channel (Red RCA connector) to GEN32 trigger in connector. - In System Plus Evolution set the Stimulator Control: Tools > Stimulator Control - Click on Start | - Stimulation with frequency set in Stimulator Control shall be displayed in Trigger Out Channel Fp2 and MKR+ channel. - GEN32 trigger in LED shall toggle at the same time Box Terminal Multi triggers out occur. - Triggers shall not display in other channels. | Passed |
| 18 | - Connect BUTTON BOX to Box Terminal Multi Trigger In 1 channel: BUTTON BOX Trigger Out channel to Box Terminal Multi Trigger In 1 channel. - Check the correct working of the button in Box Terminal Multi Trigger In 1 channel. | - Trigger in 1 channel Fp1 and MKR+ channel shall display triggers generated by the button. - Triggers shall not display in other channels. | Passed |
| 19 | - Connect BUTTON BOX to Box Terminal Multi Trigger In 2 channel: BUTTON BOX Trigger Out channel to Box Terminal Multi Trigger In 2 channel. - Check the correct working of the button in Box Terminal Multi Trigger In 2 channel. | - Trigger in 2 channel Fpz and MKR+ channel shall display triggers generated by the button. - Triggers shall not display in other channels. | Passed |
| 20 | - Start EEG acquisition (without recording) using 32 channels and 16kHz sample rate. - Leave acquisition running for 3h. | - There shall be no problems with Box Terminal Multi | Passed |

1748. Box Terminal Multi Regression - 64 Channels Acquisition Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Connect Box Terminal Multi and SD LTM 64 PLUS with HBC LTM PLUS cable. | - Box Terminal Multi display shall turn on and the headbox icon shall become green. | Passed |
| 2 | - Connect waveform generator to SD LTM 64 PLUS: - Connect white MMLTM jackbox to SD LTM 64 PLUS. - Connect MMLTM jackbox to MMADP with appropriate cables into related connectors. - MMDAP must have the switch in 1/1000 mode. |  | Passed |
| 3 | - Set in the waveform generator a sinusoidal wave Vpp=100mV f=10Hz. |  | Passed |
| 4 | - Open SystemPLUS Evolution and start an EEG recording using 64 channels and 256Hz sampling rate. | - System Plus Evolution shall display the same sinusoidal waveform in each channel. | Passed |
| 5 | - Disconnect SD LTM 64 PLUS and connect it again after a while to check the correct data recovery. | - The displayed traces shall freeze until, after few second, the recovery starts | Passed |
| 6 | - Switch Box Terminal Multi off during acquisition. - Check the correct data recovery after turning Box Terminal Multi on again. | - The displayed traces shall freeze until, after Box Terminal Multi boot, the recovery starts. | Passed |
| 7 | - Stop the acquisition. | - Acquisition shall stop. - On Box Terminal Multi display, PC icon shall switch red. | Passed |
| 8 | - Start EEG recording using 64 channels and 16384Hz sampling rate | - SystemPLUS Evolution shall display the same sinusoidal waveform in each channel. | Passed |
| 9 | - Disconnect SD LTM 64 PLUS and connect it again to check the correct data recovery. | - The displayed traces shall freeze until the recovery starts. | Passed |
| 10 | - Switch Box Terminal Multi off during acquisition. - Check the correct data recovery after turning Box Terminal Multi on again | - The displayed traces shall freeze until, after Box Terminal Multi boots, the recovery starts. | Passed |
| 11 | - Stop the acquisition. | - Acquisition shall stop. - On Box Terminal Multi display, PC icon shall switch red. | Passed |
| 12 | - Check the correct working of triggers IN: - In System Plus Evolution set triggers: Tools > Settings > Trigger Settings - Set: - Trigger in on channel Fp1 - Trigger in 2 on channel Fpz - Trigger OUT on channel Fp2 - Click on Exit - Keep only Fp1, Fpz, Fp2 and MKR+ data traces. - With a RCA cable, connect GEN32 trigger out RCA connector to BOX TERMINAL MULTI trigger in 1 channel (Black RCA connector). - Switch on GEN32 - Start Acquisition at 16384Hz (See Trigger_Settings attached) | - Trigger in 1 channel Fp1 and MKR+ channel shall display triggers at the frequency set in GEN32. - Triggers shall not display in other channels. | Passed |
| 13 | - Set time base 1 second to check if multiple triggers occur. - Keep this setting for next trigger check steps. | - Single trigger signal shall be displayed. | Passed |
| 14 | - Check trigger in 2 channels: - Disconnect RCA cable from Box Terminal Multi trigger in 1 channel (Black RCA connector) and connect it to Box Terminal Multi trigger in 2 channels (White RCA connector). | - Trigger in 2 channel Fpz and MKR+ channel shall display triggers at the frequency set in GEN32. - Triggers shall not display in other channels. | Passed |
| 15 | - Check if multiple triggers occur. | - Single trigger signal should be displayed. | Passed |
| 16 | - Verify the correct working of triggers OUT: - Connect GEN32 trigger out RCA connector to Box Terminal Multi trigger out channel (Red RCA connector). | - Trigger Out Channel shall not display any triggers. | Passed |
| 17 | - Verify the correct working of triggers OUT: - With RCA cable, connect Box Terminal Multi Trigger Out channel (Red RCA connector) to GEN32 trigger in connector. - In System Plus Evolution set the Stimulator Control: Tools > Stimulator Control - Click on Start | - Stimulation with frequency set in Stimulator Control shall be displayed in Trigger Out Channel Fp2 and MKR+ channel. - GEN32 trigger in LED shall toggle at the same time Box Terminal Multi triggers out occur. - Triggers shall not display in other channels. | Passed |
| 18 | - Connect BUTTON BOX to Box Terminal Multi Trigger In 1 channel: BUTTON BOX Trigger Out channel to Box Terminal Multi Trigger In 1 channel. - Check the correct working of the button in Box Terminal Multi Trigger In 1 channel. | - Trigger in 1 channel Fp1 and MKR+ channel shall display triggers generated by the button. - Triggers shall not display in other channels. | Passed |
| 19 | - Connect BUTTON BOX to Box Terminal Multi Trigger In 2 channel: BUTTON BOX Trigger Out channel to Box Terminal Multi Trigger In 2 channel. - Check the correct working of the button in Box Terminal Multi Trigger In 2 channel. | - Trigger in 2 channel Fpz and MKR+ channel shall display triggers generated by the button. - Triggers shall not display in other channels. | Passed |
| 20 | - Start EEG acquisition (without recording) using 64 channels and 16kHz sample rate. - Leave acquisition running for 3h. | - There shall be no problems with Box Terminal Multi | Passed |

1749. Box Terminal Multi Regression - 128 Channels Acquisition Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Set the two headboxes acquisition order: - TO ACQUIRE > MODE > DEVICE and with left-right arrows set first and second. - Then return to acquire menu. |  | Passed |
| 2 | - Connect Box Terminal Multi and SD LTM 128 PLUS with HBC LTM PLUS cable. | - Box Terminal Multi display shall turn on and the headbox icon shall become green. | Passed |
| 3 | - Connect waveform generator to SD LTM 128 PLUS: - Connect white MMLTM jackbox to SD LTM 128 PLUS. - Connect MMLTM jackbox to MMADP with appropriate cables into related connectors. - MMDAP must have the switch in 1/1000 mode. |  | Passed |
| 4 | - Set in the waveform generator a sinusoidal wave Vpp=100mV f=10Hz. |  | Passed |
| 5 | - Open SystemPLUS Evolution and start an EEG recording using 128 channels and 256Hz sampling rate. | - System Plus Evolution shall display the same sinusoidal waveform in each channel. | Passed |
| 6 | - Disconnect SD LTM 128 PLUS and connect it again after a while to check the correct data recovery. | - The displayed traces shall freeze until, after few second, the recovery starts | Passed |
| 7 | - Switch Box Terminal Multi off during acquisition. - Check the correct data recovery after turning Box Terminal Multi on again. | - The displayed traces shall freeze until, after Box Terminal Multi boot, the recovery starts. | Passed |
| 8 | - Stop the acquisition. | - Acquisition shall stop. - On Box Terminal Multi display, PC icon shall switch red. | Passed |
| 9 | - Start EEG recording using 128 channels and 8192Hz sampling rate | - SystemPLUS Evolution shall display the same sinusoidal waveform in each channel. | Passed |
| 10 | - Disconnect SD LTM 128 PLUS and connect it again to check the correct data recovery. | - The displayed traces shall freeze until the recovery starts. | Passed |
| 11 | - Switch Box Terminal Multi off during acquisition. - Check the correct data recovery after turning Box Terminal Multi on again | - The displayed traces shall freeze until, after Box Terminal Multi boots, the recovery starts. | Passed |
| 12 | - Stop the acquisition. | - Acquisition shall stop. - On Box Terminal Multi display, PC icon shall switch red. | Passed |
| 13 | - Check the correct working of triggers IN: - In System Plus Evolution set triggers: Tools > Settings > Trigger Settings - Set: - Trigger in on channel Fp1 - Trigger in 2 on channel Fpz - Trigger OUT on channel Fp2 - Click on Exit - Keep only Fp1, Fpz, Fp2 and MKR+ data traces. - With a RCA cable, connect GEN32 trigger out RCA connector to BOX TERMINAL MULTI trigger in 1 channel (Black RCA connector). - Switch on GEN32 - Start Acquisition at 8192Hz (See Trigger_Settings attached) | - Trigger in 1 channel Fp1 and MKR+ channel shall display triggers at the frequency set in GEN32. - Triggers shall not display in other channels. | Passed |
| 14 | - Set time base 1 second to check if multiple triggers occur. - Keep this setting for next trigger check steps. | - Single trigger signal shall be displayed. | Passed |
| 15 | - Check trigger in 2 channels: - Disconnect RCA cable from Box Terminal Multi trigger in 1 channel (Black RCA connector) and connect it to Box Terminal Multi trigger in 2 channels (White RCA connector). | - Trigger in 2 channel Fpz and MKR+ channel shall display triggers at the frequency set in GEN32. - Triggers shall not display in other channels. | Passed |
| 16 | - Check if multiple triggers occur. | - Single trigger signal should be displayed. | Passed |
| 17 | - Verify the correct working of triggers OUT: - Connect GEN32 trigger out RCA connector to Box Terminal Multi trigger out channel (Red RCA connector). | - Trigger Out Channel shall not display any triggers. | Passed |
| 18 | - Verify the correct working of triggers OUT: - With RCA cable, connect Box Terminal Multi Trigger Out channel (Red RCA connector) to GEN32 trigger in connector. - In System Plus Evolution set the Stimulator Control: Tools > Stimulator Control - Click on Start | - Stimulation with frequency set in Stimulator Control shall be displayed in Trigger Out Channel Fp2 and MKR+ channel. - GEN32 trigger in LED shall toggle at the same time Box Terminal Multi triggers out occur. - Triggers shall not display in other channels. | Passed |
| 19 | - Check if multiple triggers occur | - Single trigger signal shall be displayed. | Passed |
| 20 | - Connect BUTTON BOX to Box Terminal Multi Trigger In 1 channel: BUTTON BOX Trigger Out channel to Box Terminal Multi Trigger In 1 channel. - Check the correct working of the button in Box Terminal Multi Trigger In 1 channel. | - Trigger in 1 channel Fp1 and MKR+ channel shall display triggers generated by the button. - Triggers shall not display in other channels. | Passed |
| 21 | - Connect BUTTON BOX to Box Terminal Multi Trigger In 2 channel: BUTTON BOX Trigger Out channel to Box Terminal Multi Trigger In 2 channel. - Check the correct working of the button in Box Terminal Multi Trigger In 2 channel. | - Trigger in 2 channel Fpz and MKR+ channel shall display triggers generated by the button. - Triggers shall not display in other channels. | Passed |
| 22 | - Start EEG acquisition (without recording) using 128 channels and 8kHz sample rate. - Leave acquisition running for 3h. | - There shall be no problems with Box Terminal Multi | Passed |

1750. Box Terminal Multi Regression - 256 Channels Acquisition Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Set the four headboxes acquisition order: - TO ACQUIRE > MODE > DEVICE and with left-right arrows set first and second. - Then return to acquire menu. |  | Passed |
| 2 | - Connect Box Terminal Multi and SD LTM 256 PLUS with HBC LTM PLUS cable. | - Box Terminal Multi display shall turn on and the headbox icon shall become green. | Passed |
| 3 | - Connect waveform generator to SD LTM 256 PLUS: - Connect white MMLTM jackbox to SD LTM 256 PLUS. - Connect MMLTM jackbox to MMADP with appropriate cables into related connectors. - MMDAP must have the switch in 1/1000 mode. |  | Passed |
| 4 | - Set in the waveform generator a sinusoidal wave Vpp=100mV f=10Hz. |  | Passed |
| 5 | - Open SystemPLUS Evolution and start an EEG recording using 256 channels and 256Hz sampling rate. | - System Plus Evolution shall display the same sinusoidal waveform in each channel. | Passed |
| 6 | - Disconnect SD LTM 256 PLUS and connect it again after a while to check the correct data recovery. | - The displayed traces shall freeze until, after few second, the recovery starts | Passed |
| 7 | - Switch Box Terminal Multi off during acquisition. - Check the correct data recovery after turning Box Terminal Multi on again. | - The displayed traces shall freeze until, after Box Terminal Multi boot, the recovery starts. | Passed |
| 8 | - Stop the acquisition. | - Acquisition shall stop. - On Box Terminal Multi display, PC icon shall switch red. | Passed |
| 9 | - Start EEG recording using 256 channels and 16384Hz sampling rate | - SystemPLUS Evolution shall display the same sinusoidal waveform in each channel. | Passed |
| 10 | - Disconnect SD LTM 256 PLUS and connect it again to check the correct data recovery. | - The displayed traces shall freeze until the recovery starts. | Passed |
| 11 | - Switch Box Terminal Multi off during acquisition. - Check the correct data recovery after turning Box Terminal Multi on again | - The displayed traces shall freeze until, after Box Terminal Multi boots, the recovery starts. | Passed |
| 12 | - Stop the acquisition. | - Acquisition shall stop. - On Box Terminal Multi display, PC icon shall switch red. | Passed |
| 13 | - Check the correct working of triggers IN: - In System Plus Evolution set triggers: Tools > Settings > Trigger Settings - Set: - Trigger in on channel Fp1 - Trigger in 2 on channel Fpz - Trigger OUT on channel Fp2 - Click on Exit - Keep only Fp1, Fpz, Fp2 and MKR+ data traces. - With a RCA cable, connect GEN32 trigger out RCA connector to BOX TERMINAL MULTI trigger in 1 channel (Black RCA connector). - Switch on GEN32 - Start Acquisition at 16384Hz (See Trigger_Settings attached) | - Trigger in 1 channel Fp1 and MKR+ channel shall display triggers at the frequency set in GEN32. - Triggers shall not display in other channels. | Passed |
| 14 | - Set time base 1 second to check if multiple triggers occur. - Keep this setting for next trigger check steps. | - Single trigger signal shall be displayed. | Passed |
| 15 | - Check trigger in 2 channels: - Disconnect RCA cable from Box Terminal Multi trigger in 1 channel (Black RCA connector) and connect it to Box Terminal Multi trigger in 2 channels (White RCA connector). | - Trigger in 2 channel Fpz and MKR+ channel shall display triggers at the frequency set in GEN32. - Triggers shall not display in other channels. | Passed |
| 16 | - Check if multiple triggers occur. | - Single trigger signal should be displayed. | Passed |
| 17 | - Verify the correct working of triggers OUT: - Connect GEN32 trigger out RCA connector to Box Terminal Multi trigger out channel (Red RCA connector). | - Trigger Out Channel shall not display any triggers. | Passed |
| 18 | - Verify the correct working of triggers OUT: - With RCA cable, connect Box Terminal Multi Trigger Out channel (Red RCA connector) to GEN32 trigger in connector. - In System Plus Evolution set the Stimulator Control: Tools > Stimulator Control - Click on Start | - Stimulation with frequency set in Stimulator Control shall be displayed in Trigger Out Channel Fp2 and MKR+ channel. - GEN32 trigger in LED shall toggle at the same time Box Terminal Multi triggers out occur. - Triggers shall not display in other channels. | Passed |
| 19 | - Check if multiple triggers occur | - Single trigger signal shall be displayed. | Passed |
| 20 | - Connect BUTTON BOX to Box Terminal Multi Trigger In 1 channel: BUTTON BOX Trigger Out channel to Box Terminal Multi Trigger In 1 channel. - Check the correct working of the button in Box Terminal Multi Trigger In 1 channel. | - Trigger in 1 channel Fp1 and MKR+ channel shall display triggers generated by the button. - Triggers shall not display in other channels. | Passed |
| 21 | - Connect BUTTON BOX to Box Terminal Multi Trigger In 2 channel: BUTTON BOX Trigger Out channel to Box Terminal Multi Trigger In 2 channel. - Check the correct working of the button in Box Terminal Multi Trigger In 2 channel. | - Trigger in 2 channel Fpz and MKR+ channel shall display triggers generated by the button. - Triggers shall not display in other channels. | Passed |
| 22 | - Start EEG acquisition (without recording) using 256 channels and 8kHz sample rate. - Leave acquisition running for 3h. | - There shall be no problems with Box Terminal Multi | Passed |

1751. Box Terminal Multi Regression - Camera Control Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Set the acquisition system with two Box Terminal Multis (updated at latest version), two IP cameras, one Video Mixer, one SD LTM PLUS, one acquisition PC. |  | Passed |
| 2 | - Launch SpySysPlus to control whether communication errors occur. | - SpySysPlus icon shall pop up in Windows notification area. | Passed |
| 3 | - Click on SpySysPlus icon in Windows notification area. | - SpySysPlus window shall appear | Passed |
| 4 | - Enable the SpySysPlus Log viewer clicking on second icon (SpyPlus_1 attached) | - SpySysPlus Log viewer shall enable (SpyPlus_2 attached) | Passed |
| 5 | - Launch SystemPLUS Evolution | - SystemPLUS Evolution shall correctly open | Passed |
| 6 | - Set the amplifier in Tools > Options... > EEG: - Select the Headbox menu. Below Headbox set BQBroadCast in the specific dropdown menu. - Click on Set Device and select the SD LMT PLUS. | - EEG Options window shall close. | Passed |
| 7 | - ​​​​​​In the EEG Option window, select the Video Source menu. - Set the specific Video Acquisition Input Type and enable Use external camera control (over IP). - Click on Exit to close the EEG Options. | - EEG Options window shall close. | Passed |
| 8 | - Start a new acquisition - Enable Video. | - Video Acquisition window shall display. - Camera Control Panel shall display. | Passed |
| 9 | - Click on Camera Control setting icon in the top right angle. | - Camera Control setup window shall display. | Passed |
| 10 | - Set-up the two cameras, the Video Mixer and the two Box Terminal Multis: type the IP address of the devices in the specific fields. - Click OK | - Camera Control setup window shall close. | Passed |
| 11 | - Connect the amplifier to the first Box Terminal Multi and start acquisition at 256Hz. | - Acquisition shall start. | Passed |
| 12 | - Check SpyPlus Log viewer displays the correct messages. | - Log Viewer shall display broadcast received and then Start acquisition on LTM messages. | Passed |
| 13 | - Check the camera is linked with the specific Box Terminal Multi | - Camera Control window shall display the Box Terminal Multi location set-up previously. | Passed |
| 14 | - Disconnect the amplifier from the first Box Terminal Multi and connect it to the second. | - Acquisition shall freeze and shall restart immediately after a few seconds recovery. | Passed |
| 15 | - Check that SpyPlus Log viewer displays that the broadcast is received and the acquisition restart correctly. | - Log Viewer shall display broadcast received message and then shall display that the acquisition restarted with the correct recovered data. | Passed |
| 16 | - Check the camera is linked with the specific Box Terminal Multi | - The video acquisition shall switch to the second camera. Camera Control window shall display the second Box Terminal Multi location set-up previously. | Passed |

1752. Box Terminal Multi Regression – Synchronism with IP HD Camera Test Case

| Step | Description | Expected Result | Result |
| --- | --- | --- | --- |
| 1 | - Start SystemPLUS Evolution and set the recording environment: - SD LTM 256 PLUS - IP HD camera (Axis) - Video Sync to test synchronism - Set sampling frequency to 8192 Hz |  | Passed |
| 2 | - Start a Video EEG recording |  | Passed |
| 3 | - Let the recording run for >72h hours. | - Let the recording run for >72h hours. | Passed |
| 4 | - Check there are no problems with Box Terminal Multi | - Box Terminal Multi shall be still running. | Passed |
| 5 | - In review, check the correct acquisition and EEG/video/audio synchronism. | - Shift between EEG/video/audio shall be maximum +-3 frames | Passed |